## Codex Wave 3
Status | Priority | Recommended Level | Tool Family | Recommended Reference | Supplier | Verified Product Link | Unit Price (€ ex VAT) | Calibration / Verification | Calib. Period | Lifespan | Calib. Cost (€/yr) | Source Basis | Implementation Comment
Ready to integrate | Critical | Technician | Electrical safety | Fluke T150 two-pole voltage and continuity tester | Fluke | https://www.fluke.com/en/product/electrical-testing/basic-testers/fluke-t90-t110-t130-t150 | 191 | Yes | Annual | 10 years | 76.5 | Fluke France price list lists FLUKE-T150 at 191 EUR HT; official Fluke product page confirms IEC EN 61243-3 absence-of-voltage use. | Best fit to close the missing PSD technician VAT gap.
Ready to integrate | High | Technician | Wire preparation | JOKARI SECURA 2K 20100 automatic wire stripper 0.2–6.0 mm² | JOKARI | https://jokari.de/en/SECURA-2K-2.htm | 38.66 | No | Annual visual / blade condition | 5 years |  | Official JOKARI page confirms 0.2–6.0 mm²; official RRP gross is 46 EUR, converted here to approx. 38.66 EUR ex VAT. | Good field choice for sensor, command and auxiliary wire repairs.
Candidate – scope dependent | High | Technician or Team | Ferrule crimping | Weidmuller PZ 6 Roto L ferrule crimping tool | Weidmuller | https://www.weidmuller.com/en/products/workplace_accessories/tools/crimping.jsp | 145 | No | Annual inspection | 7 years |  | Weidmuller official crimping page recommends PZ 6 Roto L; current distributor pricing seen around 145 EUR in Feltron-Zeissler catalogue / higher at some resellers. | Integrate only if PSD maintenance scope includes ferrule termination in the field.
Ready to integrate | Medium | Technician or Team | Identification / labelling | Brother PT-E310BTVP industrial label printer | Brother | https://store.brother.fr/appareils/imprimantes-d-etiquettes/p-touch/pt/pte310btvp | 189 | No | Battery / cutter check | 5 years |  | Brother France product page shows 189 EUR HT and confirms electrical labelling functions. | Useful if PSD maintenance includes cable, terminal and cabinet relabelling.
Ready to integrate | High | Team | ESD cleaning | Kärcher NT 30/1 Tact Te M antistatic safety vacuum | Kärcher | https://www.kaercher.com/it/professionale/aspiratori/aspiratori-solidi-liquidi/classe-tact/nt-30-1-tact-te-m-11482350.html | 702.52 | No | Annual electrical / filter inspection | 7 years |  | Kärcher official page confirms conductive antistatic system; Klium current price 702.52 EUR ex VAT. | Recommended primary replacement for compressed-air-only cleaning logic.
Ready to integrate | Medium | Team | Safe access | Tubesca-Comabi SHERPAMATIC FIXE 8 marches (02272158) | Tubesca-Comabi | https://www.tubesca-comabi.com/fr/sherpamatic-fixe | 808.99 | No | Annual inspection | 10 years | 35.00 (est.) | Tubesca official page confirms 8-step model / max working height about 3.88–3.89 m; Master Outillage lists ref 02272158 at 808.99 EUR HT. | Good candidate for PSD header and upper equipment access in station environment.
Ready to integrate | Medium | Team | Portable ESD field kit | RS PRO ESD Field Kit 123-5615 | RS PRO | https://uk.rs-online.com/web/p/esd-grounding/1235615 | 109.6 | No | Annual resistance check | 5 years |  | RS UK lists 95.53 GBP ex VAT; converted here to approx. 109.62 EUR using March 2026 FX. | More coherent team asset than keeping full-size bench ESD equipment in the field kit.
Decision with existing rows | Critical | Team vs Depot | Closing force measurement | Choose one standard platform: keep BIA-600 BT or keep BIA-1 Class 1 | PCE / DriveTest | https://www.pce-instruments.com/us/measuring-instruments/test-meters/closing-force-measuring-device-closing-force-transducer-pce-instruments-closing-force-transducer-bia-1-bt-for-bus-and-train-doors-and-sliding-steps-det_6126330.htm | n/a | Yes | Annual ISO calibration | 10 years | n/a | Both existing rows are valid families, but overlapping. | Recommendation: standardise on one platform before simulator integration.
Reclassify existing rows | Medium | Depot | ESD workstation | Move existing Team #2 and Team #3 to Depot as main bench ESD setup | RS Components | Existing rows in workbook | Existing workbook prices | No | Annual resistance test | 5 years |  | Current workbook already contains suitable refs for mat and wrist straps. | Keep only a portable field kit at team level.
Ready to integrate | High | Depot | Fume extraction | Weller Zero Smog EL Kit 1 | Weller | https://www.weller-tools.com/eu/de/filtration/produkte/loetdaempfe/zero-smog-el-kit-1 | 855 | No | Annual airflow / filter service | 7 years |  | Weller official EU page lists Zero Smog EL Kit 1 at 855 EUR ex VAT. | Should be added together with the soldering station, not later.
Ready to integrate | High | Depot | Bench diagnostics | Keysight 34461A 6.5-digit bench DMM | Keysight | https://www.datatec.eu/de/en/keysight-34461a-tisch-system-dmm | 1363 | Yes | Annual | 10 years | 140.00 (est.) | Datatec lists 1,363 EUR ex VAT and includes calibration certificate in delivery. | Appropriate depot electronics-bench reference meter.
Ready to integrate | High | Depot | Bench power supply | Rigol DP832A programmable laboratory DC power supply | Rigol | https://rigolshop.eu/products/power-supply/dp800/dp832a.html | 449.1 | Yes | Annual performance verification | 8 years | 90.00 (est.) | Rigol official EU store lists 449.10 EUR ex VAT. | Suitable depot PSU for controlled PCB and low-voltage equipment tests.
Ready to integrate | Medium | Depot | Inspection microscope | KERN OZP 556 stereo zoom microscope | KERN | https://it.rs-online.com/web/p/microscopi/1363512 | 1591.66 | No | Annual inspection | 10 years |  | RS Italy lists 1,591.66 EUR ex VAT for OZP 556. | Strong fit for board inspection, connector checks and solder-joint review.
Ready to integrate | Medium | Depot | PCB handling | Weller Xcelite T0051502599N / Euro Solder Fix 120 PCB holder | Weller Xcelite | https://befr.rs-online.com/web/p/pcb-holders/2600911 | 101.97 | No | Replace when worn | 10 years |  | RS France lists 101.97 EUR indicative; multiple EU sellers also around 79–102 EUR. | Useful accessory to complete the depot rework bench around the soldering station.
Candidate – still needs exact EU source | Medium | Depot | ESD workstation monitor | SCS 724 workstation monitor | SCS | https://www.hisco.com/Brands/SCS/Catalog/Static-Control/ESD-Test-Measurement/Workstation-Monitors/724-15078 | 225.49 | No | Functional verification | 7 years |  | US distributor price 263.60 USD converted to approx. 225.49 EUR; EU power-adapter version still needs a cleaner local source. | Keep as a phase-4 bench maturity item if depot board work is frequent.
Pending rebuild | Critical | Team or Depot | Encoder diagnostics | Replace ambiguous Team #15 with one exact encoder diagnostic product | TBD | TBD | TBD | TBD | TBD | TBD | Current row is not reliable enough to source from. | Needs scope clarification: simple signal check, encoder simulation, or full drive diagnostics?

## Codex Wave 2
Priority | Recommended Level | Action Type | Current Sheet / No. | Proposed Tool or Decision | Operational Reason | When Used | Notes for Next Phase
Critical | Technician | Add | Technician Tools / Missing | Dedicated bipolar VAT / absence-of-voltage tester | Formal electrical safety isolation check should not rely on the Fluke 117 multimeter. | Before opening PSD cabinets or proving circuits dead. | Select one EU rail-friendly model in wave 3 and source exact price/link.
High | Technician | Add | Technician Tools / Missing | Control wiring wire stripper (0.2–6 mm²) | Small PSD corrective repairs often require clean stripping of sensor, command and auxiliary wiring. | Connector repair, limit switch replacement, harness repair. | Could be individual or shared depending maintenance scope.
High | Technician | Add | Technician Tools / Missing | Ferrule crimping pliers for bootlace ferrules | Useful if technicians terminate or reterminate PSD control wiring in cabinets or junction boxes. | Terminal replacement, cabinet rewiring, small harness rework. | If rewiring is only done by teams, move this recommendation to Team.
Medium | Technician | Add | Technician Tools / Missing | Compact label / wire marker solution | Improves traceability after cable, fuse or terminal changes in PSD cabinets. | Any intervention that changes identification or terminal numbering. | Could be 1 per team instead if technician scope is narrower.
High | Team | Add | Team Tools / Missing | ESD-safe vacuum cleaner for cabinets and PCB cleaning | Safer and more defensible than compressed-air-only cleaning; removes dust instead of redistributing it. | Preventive cleaning of DCU / PSC boards, cabinets, encoder areas. | Priority replacement candidate for Team Tools #24 logic.
Medium | Team | Add | Team Tools / Missing | Portable non-conductive step ladder / access platform | PSD header access often requires safe elevated work inside stations and technical zones. | Header beam, operator cassette, cable route and sensor access. | Could be shared across several teams depending site organisation.
Medium | Team | Add | Team Tools / Missing | Portable ESD field intervention kit | If board swaps are done on platform, a compact foldable ESD kit is more realistic than relying only on depot bench equipment. | On-platform PCB replacement or emergency electronics intervention. | Could reuse current mat / wrist strap logic in a more portable form.
Medium | Team | Add | Team Tools / Missing | Cable termination / ferrule tool set | Team-level wiring repairs likely need a more robust preparation and crimping kit than basic pliers alone. | Harness replacement, motor/encoder/sensor rewiring. | Define exact cross-section range in wave 3.
Critical | Team | Rationalise | Team Tools / #4 and #9 | Choose one standard PSD closing-force measurement platform | Two overlapping systems add cost, calibration and training complexity. | Annual compliance, post-repair verification, commissioning checks. | Keep only one unless there is a deliberate field-vs-reference split.
Medium | Team | Reclassify | Team Tools / #2 and #3 | Move main ESD mat and wrist-strap concept to Depot, keep only portable field version for team | Full-size ESD bench equipment is more coherent at depot than in routine field kit. | Workshop board handling and advanced electronics work. | Team can keep a lightweight field ESD kit if needed.
Medium | Team | Reclassify | Team Tools / #14 | Review whether battery analyser should be Team or Depot shared asset | UPS battery campaigns are periodic, not daily team work. | Scheduled battery health campaigns. | May be more economical as 1 per depot or 1 per project.
Medium | Depot | Add | Depot Tools / Missing | ESD bench mat + wrist straps + grounding point monitor | Depot board repair should start from a complete ESD workstation, not only a soldering station. | All PCB handling and bench diagnostics. | Can absorb the current Team #2 and #3 concept.
High | Depot | Add | Depot Tools / Missing | Fume extraction for soldering / rework bench | Mandatory complement to the soldering station for safe electronics workshop use. | Board repair and soldering/rework operations. | Should be added with the same priority as the Weller station.
High | Depot | Add | Depot Tools / Missing | Bench DC power supply | Needed for powered board checks, low-voltage simulation and controlled energisation outside the door system. | PCB diagnostics, sensor checks, actuator bench tests. | Define required channels and voltage ranges in wave 3.
High | Depot | Add | Depot Tools / Missing | Inspection magnifier / stereo microscope | Strongly improves PCB fault finding, connector inspection and solder joint quality control. | Component inspection and board-level repair. | Very common workshop asset for PSD electronics maintenance.
Medium | Depot | Add | Depot Tools / Missing | Bench DMM / electronics diagnostics instrument | A depot electronics bench should not rely only on field multimeters. | Component-level diagnostics and controlled measurements. | Can be simple or advanced depending repair scope.
Medium | Depot | Add | Depot Tools / Missing | Board holder / rework accessories / nozzles / consumables | Workshop effectiveness depends on the surrounding ecosystem, not only the main soldering station. | Any soldering, rework or board repair task. | Group as one depot tooling family if preferred.
Medium | Depot | Add | Depot Tools / Missing | Antistatic cleaning station or vacuum for workshop electronics | Supports safe board cleaning before and after repair. | PCB cleaning and final inspection. | Could be common with the team-level ESD vacuum family if standardised.
Medium | All levels | Scope review | Whole workbook | Define explicit maintenance scope boundaries | Some missing/extra tools depend on whether technicians do only module swap or also cable termination and board repair. | Before final price and quantity freeze. | This scope decision should drive wave 3 and simulator integration.

## Codex Review
Priority | Sheet | No. | Status | Issue | Recommended Action | Source / Basis
Critical | Technician Tools | 1 | Corrected in place | Fluke 117 was described as formal absence-of-voltage safety tool. | Keep as diagnostic multimeter and add a dedicated bipolar VAT to PSD technician list. | Fluke 117 product scope; maintenance safety practice.
Critical | Technician Tools | 5 | Corrected in place | Actik Core / PIXA 3 / IECEx were mixed in one line. | Normalised to Petzl ACTIK CORE general inspection lamp. Add ATEX lamp only if PSD hazard study requires it. | Petzl ACTIK CORE official page.
Critical | Technician Tools | 13 | Corrected in place | Old Knipex 00 19 72 LE was a belt pouch, not a shoulder tool bag. | Replaced with Knipex 00 21 02 LE Tool Bag Classic. | Knipex official product page and 2025/26 price list.
Critical | Technician Tools | 18 | Corrected in place | Old line mixed Fluke 324 and Fluke 376 FC. | Normalised around Fluke 376 FC. Consider moving this advanced clamp meter to Team level. | Fluke 376 FC sources.
Critical | Team Tools | 4 and 9 | Needs decision | Two different PSD closing-force systems likely overlap. | Select one standard platform unless depot-vs-field split is intentionally required. | PCE / DriveTest BIA product lines.
Critical | Team Tools | 15 | Needs rebuild | Encoder tester line mixes Wachendorff and Metrel with no single verified product. | Replace with one concrete encoder diagnostic product or move to depot if truly specialist. | Web verification could not confirm a clean match.
Critical | Depot Tools | All | Needs expansion | Depot sheet currently contains only one soldering station. | Build a full depot bench list: ESD workstation, fume extraction, microscope, bench PSU, test bench instrumentation, rework accessories. | Workshop maintenance scope review.
Medium | Team Tools | 8 | Corrected in place | PCAN-USB is not a classical annually calibrated instrument. | Treat as diagnostic interface with functional verification rather than calibration. | PEAK System product nature.
Medium | Team Tools | 14 | Corrected in place | BT521 price was materially understated. | Updated to Fluke list price; review whether one per team is really needed or should be shared. | Fluke France price list.
Medium | Team Tools | 23 | Corrected in place | Fluke 123B/S price was materially understated. | Updated to Fluke list price. Verify whether 1 per team or shared specialist asset is the right allocation. | Fluke France price list.
Medium | Depot Tools | 1 | Corrected in place | WX2021 price understated and service fields inconsistent. | Price corrected; kept as non-calibrated workshop station with functional verification only. | Weller official page.
Medium | Technician Tools | 6 | Needs link refresh | Current hi-vis vest URL is a category page. | Replace with one exact vest product page when finalising sources. | RS category link only.
Medium | Team Tools | 1 | Needs link refresh | Current barrier kit URL is a category page. | Replace with exact barrier-kit product page. | RS category link only.
Medium | Team Tools | 11 and 12 | Needs link refresh | Bohle links are category pages, not exact product pages. | Replace with exact suction lifter and trolley pages. | Bohle category links only.
Medium | Team Tools | 24 | Needs strategy | Compressed-air duster can move debris deeper into PSD electronics. | Prefer an ESD-safe vacuum cleaner as the primary cleaning asset; keep duster only as secondary controlled-use consumable. | PSD electronics maintenance best practice.
Medium | Technician Tools | 2 | Corrected in place | Glove rating wording and service cost were incomplete. | Updated wording and inserted annual dielectric-test estimate. Confirm with your contracted PPE test provider later. | Salisbury class 0 norms + test-cost estimate.
Medium | Technician Tools | 4 | Corrected in place | Helmet wording mixed industrial helmet and electrical class wording. | Simplified to EN 397 industrial helmet; only use EN 50365 wording if the selected helmet actually carries it. | JSP EVO3 literature review.
Medium | Technician Tools | Whole sheet | Needs completeness review | Technician list likely still misses a formal bipolar VAT and possibly wire preparation / ferrule tooling. | Add missing safety and field wiring essentials in next review wave. | PSD maintenance task analysis.

## Technician Tools
No. | Category | Tool Description | Qty / Tech | Supplier Ref. | Supplier Name | Product Link (URL) | Unit Price (€) | Calibration | Calib. Period | Lifespan | Calib. Cost (€/yr) | Price Source | Detailed Usage
1 | Electrical Measurement | TRMS Multimeter CAT III 600V – diagnostic voltage/current/resistance/continuity + NCV | 1 | Fluke 117 | Fluke | https://www.fluke.com/fr-fr/produit/multimetres/multimetres-numeriques/fluke-117 | 349 | Yes | Annual | 10 years | 120 | Fluke France official tariff HT 2023 – art.2583647 – 349€ HT | Diagnostic checks on PSD control panels and door circuits: measure 24VDC/48VDC/110VAC supplies, continuity, resistance and live troubleshooting values. Useful for first-line fault finding, but not a substitute for a dedicated bipolar absence-of-voltage tester.
2 | Safety / PPE | Insulating gloves Class 0 (1000V AC / 1500V DC) + leather overgloves – pair | 1 | Salisbury 2750R/8H + AT127 | Honeywell Salisbury | https://www.salisburybyhoneywell.com/en/products/electrical-safety-gloves/class-0-gloves | 85 | Yes | 6 months (dielectric test) | Up to 3 years in service / replace if damaged or failed | 16 € (est.) | Honeywell Salisbury distributeur EU HT 2024 – classe 0 ~85€ | Mandatory when working on PSD control cabinets with live 24–230V circuits. Also used during UPS battery maintenance (110VDC). Store in protective pouch provided.
3 | Safety / PPE | Safety glasses anti-arc EN166 anti-scratch anti-fog | 1 | Bolle Safety Cobra COBPSI | Bolle Safety | https://www.bollesafety.com/en-int/sport/cobra | 22 | No | N/A | 1 year or if scratched |  | Bollé Safety distributeur EU HT 2024 – COBPSI ~22€ | Eye protection during all PSD interventions: grinding dust from guide rail, spring replacement, glass panel handling. Mandatory PPE per EN 166 and EN 170.
4 | Safety / PPE | Safety helmet with chin strap – industrial helmet EN 397 | 1 | JSP EVO3 REV AJF170 | JSP Safety | https://www.jspsafety.com/en-gb/products/head-protection/helmets/evo3-revolution/ | 55 | No | N/A | 3 years |  | JSP Safety distributeur EU HT 2024 – EVO3 REV ~55€ | Head protection under PSD header beams, overhead cable runs, and during glass panel extraction. Use an EN 50365-certified model only if electrical helmet insulation is specifically required by the maintenance method statement.
5 | Safety / PPE | Rechargeable headlamp 625 lm – general inspection lighting | 1 | Petzl ACTIK CORE | Petzl | https://www.petzl.com/US/en/Sport/Headlamps/ACTIK-CORE | 55,00 € | No | Annual (battery check) | 5 years |  | Petzl retail reference 2026 – ACTIK CORE list value 55€ | Hands-free lighting for PSD technical rooms, below-platform inspections, inside header cavities and during confined access. General maintenance headlamp; not an ATEX / IECEx lamp.
6 | Safety / PPE | Hi-visibility vest Class 2 EN ISO 20471 – yellow or orange | 1 | 3M 8906 Class 2 Vest | 3M / RS Components | https://uk.rs-online.com/web/c/safety-security/personal-protective-equipment/high-visibility-clothing/?searchTerm=hi-vis+vest | 18 | No | N/A | 2 years |  | RS Components EU HT 2024 – ~18€ | Mandatory on platform during service windows or near-train environments. Ensures technician visibility to drivers and station operators during PSD interventions.
7 | Safety / PPE | Personal LOTO kit – steel padlock, hasp, tags, cable – individual pouch | 1 | Brady 65674 Personal LOTO Kit | Brady | https://www.bradyid.eu/en-eu/lockout-tagout/personal-lockout-kits/65674 | 58 | No | Annual (inspection) | 10 years |  | Brady distributeur EU HT 2024 – 65674 ~58€ | Individual lockout of PSD local control boxes, UPS isolators, and 230V MCBs before any maintenance on door operators, PCBs or wiring. 1 padlock per technician mandatory.
8 | Safety / PPE | Lockout padlock steel – 1 unique key – red – S410RED | 2 | Master Lock S410RED | Master Lock | https://www.masterlock.eu/products/padlocks/safety-lockout-padlocks/S410RED | 16 | No | 5 years | 5 years |  | Master Lock EU HT 2024 – S410RED ~16€ | 1 key = 1 technician. Carry 2 padlocks: 1 in use + 1 spare. Applied to LCB (Local Control Box) before opening door operator enclosure.
9 | General Hand Tools | VDE 1000V insulated screwdriver set – 7 pcs – SoftFinish | 1 | Wiha 36295 SoftFinish VDE | Wiha | https://www.wiha.com/eu/en/products/screwdrivers/vde-screwdrivers/softfinish-vde-set-in-a-roll/36295/ | 78 | No | Annual (visual) | 10 years |  | Wiha direct EU HT 2024 – 36295 ~78€ | Opening PSD Local Control Board terminal covers, replacing fuses, tightening PCB connectors, adjusting limit switch screws, securing door panel fixings.
10 | General Hand Tools | VDE 1000V insulated pliers set – 5 pcs – 002012 | 1 | Knipex 00 20 12 VDE | Knipex | https://www.knipex.com/en/products/002012 | 145 | No | Annual (visual) | 10 years |  | Knipex direct EU HT 2024 – 002012 ~145€ | Cable cutting, wire stripping and forming for PSD control wiring (0.5–6mm²), connector extraction, and terminal pin manipulation on DCU/PSC boards.
11 | General Hand Tools | Ratchet set 1/4" + 3/8" – metric sockets 4–19mm + bits – 40 pcs | 1 | Wera 9100 SB RA | Wera | https://www.wera.de/en/products/sets/ratchet-sets/kollektion-9100-sb-ra/ | 100 | No | Replace when worn | 10 years |  | Wera direct EU HT 2024 – 9100 SB RA ~100€ | Fastening and removal of door operator mounting bolts, guide rail fixings, panel frame bolts, and header beam hardware. Essential for preventive maintenance cycle.
12 | General Hand Tools | Hex key set short + long metric 1.5–10mm – 18 pcs | 1 | Wiha 352 SB 18 T-handle | Wiha | https://www.wiha.com/eu/en/products/hex-tools-and-keys/hex-keys/hex-key-set-in-a-staple-with-t-handle-18-pcs/352sb18/ | 42 | No | Replace when worn | 10 years |  | Wiha direct EU HT 2024 – 352 SB 18 ~42€ | CHC bolts on door operator cassettes, motor mounting plates, encoder brackets and alignment shims. Long-reach keys essential for recessed fasteners.
13 | General Hand Tools | Individual tool bag – classic shoulder-strap type – 11 compartments | 1 | Knipex 00 21 02 LE Tool Bag "Classic" | Knipex | https://www.knipex.com/products/tool-bags-and-tool-cases/tool-bag-classic-empty/tool-bag-classic-empty/002102LE | 132,00 € | No | Replace if worn | 7 years |  | KNIPEX price list 2025/26 – 00 21 02 LE = 132 EUR | Personal field tool bag for platform work. Carries screwdrivers, pliers, multimeter, gauges, mirror, LOTO kit and daily PSD tools while keeping both hands free during access and movement.
14 | General Hand Tools | Rubber mallet 500g – non-marking – EN 12534 | 1 | Facom BTM.500 | Facom | https://www.facom.com/fr-fr/products/hammers-and-mallets/mallets/non-sparking-mallets/btm_500/ | 22 | No | Replace if worn | 5 years |  | Facom direct EU HT 2024 – BTM.500 ~22€ | Gentle seating of door panel guides, rubber seals, threshold profiles and operator cassettes during installation or replacement. Avoids surface damage to glass frames.
15 | General Hand Tools | Feeler gauge set 20 blades 0.05–1.00mm – stainless steel – 188-101 | 1 | Mitutoyo 188-101 | Mitutoyo | https://shop.mitutoyo.eu/web/mitutoyo/en_GB/mitutoyo/01.05.01/Straight%20Feeler%20Gauge/188-101/$catalogue/mitutoyoData/PR/188-101/index.xhtml | 28 | No | Annual (visual) | 10 years |  | Mitutoyo distributeur EU HT 2024 – 188-101 ~28€ | Measurement of door-to-frame clearances, guide rail play, roller-to-track contact gaps, and lock bolt engagement depth. Critical for geometry adjustment and EN 17168 compliance.
16 | General Hand Tools | Telescopic magnetic pick-up rod – 65cm – 600g capacity | 1 | Facom DA.15PB | Facom | https://www.facom.com/fr-fr/products/lighting-and-tools-inspection/telescopic-magnetic-pick-up-tools/telescopic-magnetic-pick-up-tool/da_15pb/ | 14 | No | Replace if needed | 5 years |  | Facom direct EU HT 2024 – DA.15PB ~14€ | Retrieval of dropped screws, nuts, springs and small components inside PSD header cavities, guide rail channels and under-threshold gaps – prevents FOD (Foreign Object Debris) incidents.
17 | General Hand Tools | Telescopic inspection mirror with LED – 360° swivel | 1 | Facom S.270PB | Facom | https://www.facom.com/fr-fr/products/lighting-and-tools-inspection/inspection-mirrors/telescopic-inspection-mirror-with-light/s_270pb/ | 22 | No | Replace if needed | 5 years |  | Facom direct EU HT 2024 – S.270PB ~22€ | Visual inspection of locking bolt engagement, upper guide roller condition, door top seal wear, and hidden wiring harness routing without removing panels.
18 | Electrical Measurement | AC/DC TRMS clamp meter 1000A – CAT IV 600V – iFlex – Fluke Connect | 1 | Fluke 376 FC | Fluke | https://www.fluke.com/en-id/product/electrical-testing/clamp-meters/fluke-376-fc | 703,00 € | Yes | Annual | 10 years | 160 | Fluke price list Europe 2025/26 – FLUKE-376 FC ~703 EUR list | Current measurement on PSD operator supply feeds, UPS output circuits and control loops. The iFlex probe is useful where conductors are inaccessible. Best used as an advanced team diagnostic meter rather than a basic technician meter.
19 | Lubrication & Cleaning | Dry PTFE lubricant spray 400ml – no dust accumulation – WD-40 Specialist | 2 | WD-40 Specialist PTFE Lubricant 300013 | WD-40 / RS Components | https://uk.rs-online.com/web/p/lubricants/0382306/ | 12 | No | N/A (consumable) | N/A |  | WD-40 / RS Components EU HT 2024 – ~12€/canister | Lubrication of PSD door guide rails, rolling screw mechanism, latch pins and locking bolts per O&M manual. PTFE dry formula prevents dust and debris accumulation. Applied at each 6-month PM visit.
20 | Lubrication & Cleaning | Silicone spray 300ml – rubber seal conditioner and lubricant | 1 | Wurth Silicone Spray 0893 115 0 | Wurth | https://www.wurth.co.uk/silicon-spray | 9 | No | N/A (consumable) | N/A |  | Würth direct EU HT 2024 – 0893 115 0 ~9€ | Maintenance of PSD door rubber seals, gaskets, and flexible dust wipers. Prevents cracking, restores elasticity, reduces friction coefficient on seal-to-frame contact faces.
21 | Lubrication & Cleaning | Microfibre cloths pack 10 + glass cleaner spray 500ml | 2 | Wurth Professional Glass Cleaner + cloths | Wurth | https://www.wurth.co.uk/glass-cleaner | 14 | No | N/A (consumable) | N/A |  | Würth direct EU HT 2024 – ~14€ kit | Cleaning of tempered glass panels, polycarbonate screens, optical sensors (presence detectors, gap sensors) and laser emitter lenses. Clean optics prevent false door obstruction faults.

## Team Tools
No. | Category | Tool Description | Qty / Team | Supplier Ref. | Supplier Name | Product Link (URL) | Unit Price (€) | Calibration | Calib. Period | Lifespan | Calib. Cost (€/yr) | Price Source | Detailed Usage
1 | Safety / PPE | Safety barrier kit – retractable belt posts x4 + warning tape 10m | 2 | VISO TB250 Kit + warning tape | VISO / RS Components | https://uk.rs-online.com/web/c/safety-security/safety-barriers-barriers-and-bollards/ | 180 | No | Annual (inspection) | 7 years |  | VISO / RS Components EU HT 2024 – ~180€ | Platform zoning during PSD maintenance service windows. Establishes safe exclusion zone around work area preventing passenger intrusion during door panel removal or electrical work.
2 | Safety / PPE | Anti-static ESD ground mat 1×1m – 10^6–10^9 Ohm – EN 61340-4-1 | 2 | RS Pro ESD Mat 1m×1m 229-5231 | RS Components | https://uk.rs-online.com/web/p/esd-mats/2295231/ | 95 | No | Annual (resistance test) | 5 years |  | RS Components EU HT 2024 – ref 229-5231 ~95€ | ESD-safe work surface for PCB board handling, DCU/PSC board replacement, and control electronics maintenance. Prevents electrostatic discharge damage to servo drive electronics and communication boards.
3 | Safety / PPE | Anti-static wrist strap ESD – adjustable – 1 MOhm resistor | 4 | RS Pro ESD Wrist Strap 229-8422 | RS Components | https://uk.rs-online.com/web/p/anti-static-wrist-straps/2298422/ | 15 | No | Annual (resistance test) | 3 years |  | RS Components EU HT 2024 – ref 229-8422 ~15€/pcs | Individual grounding for each technician handling PSD control electronics (DCU boards, communication modules, servo drive PCBs). Prevents ESD damage to components costing €500–5,000 per board.
4 | PSD-Specific Measurement | Wireless door closing force system BIA-600 BT – EN 14752 – Android app + PinchPilot – ISO cert | 1 | DriveTest / PCE BIA-600 BT System (ref BIA600BT-SY-10-700) | PCE Instruments (DriveTest) | https://www.pce-instruments.com/english/measuring-instruments/test-meters/closing-force-measuring-device-pce-instruments-closing-force-measuring-device-bia-600-bt-for-rail-vehicle-doors-sliding-steps-det_6126436.htm | 2490 | Yes | Annual (ISO calibration + certificate) | 10 years | 350 | PCE/DriveTest EU HT 2024 – BIA-600 BT système complet ~2490€ | Team wireless closing force campaign testing of all doors on a platform. Bluetooth to tablet app for real-time evaluation. PinchPilot auto-generates EN 14752 compliance reports with door ID, measurement point, force value, pass/fail. QR code door identification. Mandatory for annual safety certification.
5 | PSD-Specific Measurement | Sonic belt tension meter – resonant frequency – for timing belt and V-belt drives | 1 | Gates Sonic Tension Meter 550C (ref 74200550) | Gates Corporation | https://www.gates.com/en_us/maintenance-tools/belt-maintenance/sonic-tension-meter.html | 680 | Yes | Annual | 10 years | 120 | Gates / vbeltsdirect.com distributeur EU HT 2024 – 550C ~680€ | Non-contact resonant frequency measurement of timing belt tension on PSD door operator belt drives. Correct tension per manufacturer spec (typically 100–300N) is critical – under-tension causes belt slip, over-tension causes premature bearing and belt failure. Measurement takes <10 seconds per belt. Current model replaces discontinued 507C/508C.
6 | PSD-Specific Measurement | Laser line level 3×360° self-levelling H/V/D – IP54 – range 30m | 1 | Bosch GLL 3-80 Professional (ref 0601063S00) | Bosch Professional | https://www.bosch-professional.com/gb/en/products/gll-3-80-0601063S00.html | 285 | Yes | Annual | 8 years | 95 | Bosch Professional distributeur EU HT 2024 – GLL 3-80 ~285€ | Door panel vertical and horizontal alignment verification after installation or replacement. Projects reference line along entire platform façade to verify door plumb, header beam level, and panel-to-panel step deviation. Alignment tolerance typically ±2mm. Also used for door-to-train gap survey.
7 | PSD-Specific Measurement | Thermal imaging camera 160×120px radiometric – MSX – WiFi – Fluke Connect | 1 | Fluke TiS60+ Thermal Camera | Fluke | https://www.fluke.com/fr-fr/produit/cameras-thermiques/fluke-tis60-plus | 2150 | Yes | Annual (COFRAC) | 8 years | 380 | Fluke France tarif caméras thermiques HT 2024 – TiS60+ ~2150€ | Non-contact detection of overheating door operator motors (bearing failure, winding fault), hot PCB components (failed capacitors, power MOSFETs), and connection hotspots in LCB panels. Preventive thermographic inspection of entire platform every 6 months.
8 | PSD-Specific Measurement | CAN bus / RS-485 USB analyser + PCAN-View software – 1Mbit/s | 1 | PEAK System PCAN-USB (ref IPEH-002022) | PEAK System | https://www.peak-system.com/PCAN-USB.199.0.html | 175 | No | Functional check / firmware updates | 8 years |  | PEAK System official – peak-system.com HT 2024 – PCAN-USB ~175€ | Real-time monitoring and logging of PSD communication buses (CAN, RS-485 Modbus) for protocol diagnosis, message tracing and timing analysis. Treat as a diagnostic interface rather than a periodically calibrated measuring instrument.
9 | PSD-Specific Measurement | Door closing force meter BIA Class 1 – electronic – EN 14752 – with PinchPilot PC software | 1 | DriveTest / PCE BIA-1 Original Class 1 + SEB2 kit | PCE Instruments (DriveTest) | https://www.pce-instruments.com/english/measuring-instruments/test-meters/closing-force-measuring-device-pce-instruments-closing-force-measuring-device-bia-class-1-det_6126442.htm | 1490 | Yes | Annual (ISO calibration certificate) | 10 years | 280 | DriveTest/PCE BIA Class 1 + SEB2 – prix concessionnaire EU HT 2024 ~1490€ | Mandatory periodic measurement of PSD closing force per EN 14752 and EN 17168. Measures peak closing force (max 150N per standard) at multiple door positions. Data logged to PinchPilot software for report generation and traceability. Used at commissioning, periodic inspection, and after any door operator or spring replacement.
10 | PSD-Specific Measurement | Digital tachometer non-contact laser + contact – 0.10 to 25,000 RPM – NIST certificate | 1 | Shimpo DT-205LR-S12 | Shimpo / Grainger | https://www.grainger.com/product/SHIMPO-Tachometer-0-006-RPM-Accuracy-21YD77 | 370 | Yes | Annual | 8 years | 95 | Grainger US 2024 – DT-205LR-S12 ~$393 / ~370€ HT | Measurement of door operator motor RPM during functional check, belt pulley speed verification, and comparison against programmed door speed profile (typically 0.4–0.6 m/s). Detects belt slippage and motor speed deviations.
11 | Glass & Panel Handling | Vacuum suction lifter 3-cup lever – 120kg – aluminium – Veribor | 2 | Bohle Veribor BO 603.021 (3-cup 120kg) | Bohle Group | https://www.bohle.com/uk-EN/online-shop/vacuum-technology/suction-lifters/ | 280 | No | Annual (seal + pressure test) | 7 years |  | Bohle Group direct EU HT 2024 – BO 603.021 ~280€ | Safe removal and installation of PSD laminated tempered glass panels (typically 20–40kg per leaf). 3-cup aluminium design provides secure hold on full-height panels. TUV GS certified. Essential for glass replacement after impact damage.
12 | Glass & Panel Handling | Glass panel transport A-frame trolley – 200kg – adjustable height | 1 | Bohle GTV 200 Glass Transport Trolley | Bohle Group | https://www.bohle.com/uk-EN/online-shop/handling/glass-transport-trolleys/ | 650 | No | Annual (inspection) | 10 years |  | Bohle Group direct EU HT 2024 – GTV 200 ~650€ | Platform-side transport of replacement glass panels from storage to installation point. A-frame maintains panels at safe angle. Accommodates full-height (2.1–3.0m) and half-height (1.2–1.5m) PSD glass panels up to 200kg.
13 | Electronics & Controls | Rugged laptop IP65 Win11 Pro – MIL-STD-810H – for PSD MCS/diagnostic software | 2 | Panasonic Toughbook FZ-55 (ref FZ-55A-009KE) | Panasonic | https://www.panasonic.com/uk/computers/toughbook/toughbook-55.html | 2800 | No | 5-year replacement | 5 years |  | Panasonic Toughbook UK HT 2024 – FZ-55 ~2800€ | Connection to PSD Monitoring & Control System (MCS/PSCD) for door status viewing, alarm history, parameter adjustment (door speed, overtorque force, opening/closing position), and firmware updates on DCU and PSC boards. Essential for Level 2+ corrective maintenance.
14 | Electronics & Controls | Stationary battery tester – internal impedance – 12–240V – IEEE 1188 / IEC 60896 | 1 | Fluke BT521 Battery Analyzer | Fluke | https://www.fluke.com/fr-fr/produit/testeurs-batteries/fluke-bt521 | 5 545,00 € | Yes | Annual | 10 years | 195 | Fluke France price list (published 2026, prices dated 1 Mar 2023) – FLUKE-BT521 = 5545 EUR HT | State-of-health testing of PSD Emergency Power Supply (EPS/UPS) sealed lead-acid batteries (typically 12V/24V strings). Internal impedance identifies capacity degradation before failure. Batteries typically require replacement at 3–5 year intervals.
15 | Electronics & Controls | Incremental encoder signal checker – A/B/Z quadrature – 5V/12V/24V | 1 | Wachendorff WDGA encoder tester / or Metrel MI3295 multifunction | RS Components / Wachendorff | https://uk.rs-online.com/web/c/test-measurement/signal-generators-siggen-accessories/ | 580 | Yes | Annual | 8 years |  | RS Components / Wachendorff EU HT 2024 – ~580€ | Verification of door operator incremental encoder outputs (A/B/Z channels, typically 100–2500 ppr) without running the full drive system. Confirms encoder integrity after motor replacement or following encoder connector damage.
16 | Workshop / Team | Rolling tool chest 3 drawers lockable – 80kg – FatMax XL | 1 | Stanley FMST1-75791 FatMax XL | Stanley Tools | https://www.stanleytools.com/product/fmst1-75791/fatmax-xl-15-drawer-mobile-combo | 245 | No | Replace when worn | 10 years |  | Stanley Tools distributeur EU HT 2024 – FMST1-75791 ~245€ | Centralised secure storage of team-level specialist tools at station or depot. Lockable drawers protect calibrated instruments (closing force meters, vibration analyser, thermal camera) and spare PCB boards.
17 | Workshop / Team | Torque wrench 3/8" 2–25 Nm – ISO 6789 Class II – with COFRAC certificate | 1 | Facom S.306-25 | Facom | https://www.facom.com/fr-fr/products/tightening-and-torque/torque-wrenches/adjustable-torque-wrenches/s_306-25/ | 155 | Yes | Annual (COFRAC) | 10 years | 75 | Facom / RS France HT 2024 – S.306-25 ~155€ | Controlled torque on PSD panel fixing bolts (M5–M8), LCB door hardware, and sensor brackets. Correct torque critical for glass panel structural integrity and vibration resistance on high-cycle door systems (500,000+ cycles/year).
18 | Workshop / Team | Torque wrench 1/2" 20–200 Nm – ISO 6789 Class II – with COFRAC certificate | 1 | Facom S.307-200 | Facom | https://www.facom.com/fr-fr/products/tightening-and-torque/torque-wrenches/adjustable-torque-wrenches/s_307-200/ | 195 | Yes | Annual (COFRAC) | 10 years | 90 | Facom / RS France HT 2024 – S.307-200 ~195€ | Higher torque fixings: PSD header beam mounting bolts (M10–M16), floor anchor bolts, and door operator cassette fixing hardware. Used for installation/replacement of full door operator units.
19 | Workshop / Team | Group LOTO station – steel hasp 13-padlock + 4 team padlocks | 2 | Brady 50931 Group Lockout Hasp + S410RED x4 | Brady / Master Lock | https://www.bradyid.eu/en-eu/lockout-tagout/group-lockout-devices/hasp-lockouts/50931 | 145 | No | Annual (inspection) | 10 years |  | Brady / Master Lock distributeur EU HT 2024 – hasp 50931 + S410RED ~145€ | Multi-technician lockout of PSD zone isolators, MCS/PSCD control panels, and 400VAC distribution boards when entire platform section requires maintenance. Each technician applies own padlock to shared hasp.
20 | Workshop / Team | Grease gun 500cc lever handle – with flexible hose – EN 1929-1 | 1 | Lincoln 1142 Standard Lever Grease Gun | Lincoln Industrial | https://www.lincolnindustrial.com/products/lubrication-tools/grease-guns/1142/ | 65 | No | Annual (inspection) | 7 years |  | Lincoln Industrial distributeur EU HT 2024 – 1142 ~65€ | Application of manufacturer-specified grease to PSD lead screw/ball screw assemblies, upper guide roller bearings, and pivot pins during scheduled preventive maintenance (typically 6-monthly or per O&M cycle count).
21 | Workshop / Team | Cordless drill/driver 18V Li-ion – 13mm chuck – 2×2Ah batteries + case | 1 | Bosch GSR 18V-55 Professional | Bosch Professional | https://www.bosch-professional.com/gb/en/products/gsr-18v-55-0615A5002N.html | 175 | No | Replace if defective | 7 years |  | Bosch Professional distributeur EU HT 2024 – GSR 18V-55 ~175€ | Removal and reinstallation of door panel fixings, guide rail brackets, header cover panels, and cable duct lids. Essential for rapid corrective maintenance during service windows.
22 | Electrical Measurement | Portable insulation resistance tester 500V/1000V DC – CAT IV 600V – Fluke 1507 | 1 | Fluke 1507 | Fluke | https://www.fluke.com/fr-fr/produit/testeurs-isolement/fluke-1507 | 699,00 € | Yes | Annual | 10 years | 145 | Fluke France price list (published 2026, prices dated 1 Mar 2023) – FLUKE-1507 = 699 EUR HT | 5-year mandatory insulation test of PSD façade wiring per O&M manual (performed with track de-energised). Also used for motor winding insulation checks and cable harness integrity before energisation.
23 | Electrical Measurement | Portable oscilloscope 2-channel 200MHz – CAT III 600V – ScopeMeter – IP51 | 1 | Fluke 123B/S Industrial ScopeMeter | Fluke | https://www.fluke.com/fr-fr/produit/oscilloscopes/fluke-123b-s-industrial-scopemeter | 2 685,00 € | Yes | Annual | 10 years | 280 | Fluke France price list (published 2026, prices dated 1 Mar 2023) – FLUKE-123B/EU/S = 2685 EUR HT | Analysis of servo motor encoder signals (A/B/Z channels), PWM drive output waveforms, door position sensor outputs (Hall effect, optical). Diagnoses intermittent faults not visible on multimeter. IP51 rated for platform environment.
24 | Lubrication & Cleaning | Compressed air duster 400ml – electronics safe – non-flammable | 2 | Kontakt Chemie Airo Duster 400ml | RS Components | https://uk.rs-online.com/web/p/air-dusters/8444510/ | 8 | No | N/A (consumable) | N/A |  | RS Components EU HT 2024 – ~8€/canister | Removal of dust, debris and metallic particles from PCB surfaces (DCU, PSC boards), encoder optical discs, and connector pins during preventive maintenance. Prevents false encoder signals.

## Depot Tools
No. | Category | Tool Description | Qty / Depot | Supplier Ref. | Supplier Name | Product Link (URL) | Unit Price (€) | Calibration | Calib. Period | Lifespan | Calib. Cost (€/yr) | Price Source | Detailed Usage
1 | Electronics & Controls | ESD soldering station dual-channel hot air + iron – 50-450°C – Weller | 1 | Weller WX2021 Dual Station | Weller Tools | https://www.weller-tools.com/professional/EUR/en/Professional/Soldering+stations/Wx-Series/WX2021.htm | 1 343,00 € | No | Functional temperature / ESD bench verification | 10 years |  | Weller official product page 2026 – WX2021 displayed at 1343 EUR HT | Workshop repair of PSD control boards (DCU, PSC, PSC-L): component-level repair of failed capacitors, fuses, relay contacts, and SMD components. Reduces cost vs full board replacement (board cost typically €400–2,500). ESD-safe bench only.